Figure S1. Egg orientation device and experimental setup. Assembled egg orientation device (EOD), showing the electrorotation module (ERM), out-of-plane mechanical rotation module, rotatable platform, stationary base, and motor control unit.

Table S1. Estimated component costs of the egg orientation device (EOD).

| Component | Manufacturer, model, or material | Estimated cost (USD) |
| --- | --- | --- |
| 3D-printed structural components | PLA filament and photopolymer resin | 10 |
| Function generator | Siglent, SDG2000X series | 574 |
| Four-output DC power supply | FLIR Extech, 382270 | 679 |
| Phase-splitter printed circuit board | Custom-built PCB and associated electronic components | 200 |
| Rotating motor | N20 Motor with encoder or stepper motor NEMA-8 | 20 |
| Total estimated component cost |  | 1483 |

| Table S2. Printing settings for 3-D Resin Printer. |  |
| --- | --- |
| Parameters | Settings |
| Printer model | CADworks3D Microfluidics 285D |
| Slicer name | Utility |
| Slicer version | 6.4.4.t12 |
| Layer height (mm) | 0.03 |
| Exposure time (s) | 0.8 |
| Base curing (s) | 25 |
| Gap adjustment (mm) | -0.25 |
| Base layer | 1 |
| Buffer layer | 4 |
| Power (%) | 100 |
| Print delay (s) | 90 |

Definitions

Layer Height: The thickness of each layer in the print.

Exposure Time: The amount of time for UV curing (seconds) per layer.

Base curing: Curing time for base layers.

Gap adjustment: Adjust the thickness of the first layer.

Base layer: Define the number of base layers.

Buffer layer: Set the Number of buffer layers.

Power: At 100% is the existing brightness of the light engine.

Print delay: For the first layer, the build plate stays for at least 1 sec. Then cure.

Printing post-processing

We placed the device in a zip lock filled with 99% isopropyl alcohol (IPA) and cleaned it in an ultrasonic water bath for 2 min to remove residual resin. The devices were subsequently rinsed with fresh 99% IPA followed by deionized water to remove remaining uncured resin. After air-drying, the devices underwent a final 1-min curing step under 405 nm UV light.

Phase splitter schematics

Figure S2. Circuit schematic of the custom four-phase signal-generation PCB. Two quadrature input signals (I and Q, separated by 90°) are split into complementary signal pairs using ADA4950-1 amplifiers and routed through four gain stages (A–D) to provide the outputs labeled 0°, 90°, 180°, and 270° for electrorotation.

Figure S3. Circuit schematic of an individual output gain stage. Each of the four channels uses an ADA4870 amplifier in a non-inverting configuration, powered by ±12 V supplies. The schematic shows the feedback resistor network, supply-decoupling capacitors, shutdown-control connection, and 4.7 Ω series resistor output. The same circuit is repeated in gain stages A–D of the four-phase signal-generation PCB.

Table S3. Bill of materials for the four-phase signal-generation PCB.

| Designator | Part Number | Value | Description | Quantity |
| --- | --- | --- | --- | --- |
| U1, U2 | ADA4950-1YCPZ-R7 | N/A | Differential amplifier | 2 |
| U3A, U3B, U3C, U3D | ADA4870ACPZ | N/A | High voltage drive amplifier | 4 |
| C1, C3, C6, C8 | C0603X102J4RECAUTO | 1nF | Power supply decoupling capacitor | 4 |
| C2, C4, C5, C7, C9, C10, C13A, C13B, C13C, C13D, C16A, C16B, C16C, C16D | CC0603JRX7R9BB104 | 100nF | Power supply decoupling capacitor, input matching | 14 |
| C11, C12 | TAJB476K010RNJ | 47uF | Bulk power supply decoupling capacitor | 2 |
| C14A, C14B, C14C, C14D, C15A, C15B, C15C, C15D | TAJB226K020RNJ | 22uF | Bulk power supply decoupling capacitor | 8 |
| J1, J2, J3, J4, J5, J6 | 73100-0105 | N/A | BNC connector | 6 |
| J7, J8 | TB002-500-03BE | N/A | Screw terminals for power supply | 2 |
| R1, R2, R3, R5, R6, R7 | TNPW060350R0DEEA | 50 Ω | Impedance matching resistor | 6 |
| R4, R8 | RP0603BRD0725RL | 25 Ω | Impedance matching resistor | 2 |
| R9 | RMCF0603FT34K8 | 35k Ω | Shutdown control | 1 |
| R10 | RT0603BRE075KL | 5k Ω | Shutdown control | 1 |
| R11A, R11B, R11C, R11D | RC0603FR-134R7L | 4.7 Ω | Output stability resistor | 4 |
| R12A, R12B, R12C, R12D | ERJ-3EKF3010V | 301 Ω | Gain setting resistor | 4 |
| R13A, R13B, R13C, R13D | RC0603FR-071K21L | 1.21 KΩ | Gain setting resistor | 4 |
| PCB | Custom PCB | N/A | Printed circuit board | 1 |